## Sheet1
Requirement ID | Section | Keyword | Requirement Text
REQ-1 |  | shall | (b)Where you have a specification or tender requirement to reproduce a document or portions of a document as part of its documentation for external submission in response to a tender, the necessary pages of the document, including the whole document if required, may be reproduced and submitted provided a copyright notice is included. You shall notify Intertek Inform of any such use. For internal and archival purposes only, a paper copy may be attached to your documentation and shall be considered a permanent part of that documentation.
REQ-2 |  | shall | This European Standard shall be given the status of a national standard, either by publication of an identical text or by endorsement, at the latest by April 2025, and conflicting national standards shall be withdrawn at the latest by April 2025.
REQ-3 |  | shall | Attention is drawn to the possibility that some of the elements of this document may be the subject of patent rights. CEN shall not be held responsible for identifying any or all such patent rights.
REQ-4 | 12.1.1 | shall | ISO draws attention to the possibility that the implementation of this document may involve the use of (a) patent(s). ISO takes no position concerning the evidence, validity or applicability of any claimed patent rights in respect thereof. As of the date of publication of this document, ISO had not received notice of (a) patent(s) which may be required to implement this document. However, implementers are cautioned that t the latest information, which may be obtained from the patent database available at . ISO shall not be held responsible for identifying any or all such patent rights.
REQ-5 | 4.1 | shall | A filter system shall be designated by its access site (see 3.3), orientation (see 3.17), implantation site (see
REQ-6 | 4.2 | shall | Materials of the filter system (e.g. wire, imaging markers, coatings) shall be described by their generic or
REQ-7 | 4.3 | shall | 4.3 Configuration and size designation for filters The configuration of a filter shall be designated by its geometry (e.g. conical) and whether it is permanent, optional or convertible. The size of a filter shall be designated by the minimum and maximum intended caval lumen diameters.
REQ-8 | 4.4 | shall | 5 Intended performance The requirements of ISO 14630:2012, Clause 4 shall apply.
REQ-9 | 6.1 | shall | 6.1 General The requirements for design attributes in 6.2 to 6.10 and of ISO 14630:2012, Clause 5 apply. General design attributes for the filter system, the filter, the filter retrieval system and the filter conversion system are listed in Tables A.3 and A.9 with reference to the nonclinical testing necessary for the evaluation of the design. It is recognized that not all tests identified in a category are necessary or practical for any given filter and/or system. The tests considered and the rationale for selection and/or waiving of tests shall be documented.
REQ-10 | 6.2 | shall | 6.2 Filter system In addition to the general requirements, the design attributes of the filter system shall at least take into account the following: a) the ability to permit safe and consistent deliverability of the filter to the intended deployment location; b) the ability to permit accurate and safe deployment of the filter; c) the ability to inject contrast via the delivery system if indicated in the IFU; d) the ability to permit safe withdrawal of the delivery system following deployment; e) the ability to maintain adequate structural integrity.
REQ-11 | 6.3 | shall | 6.3 Vena cava filter The design attributes of the vena cava filter shall at least take into account the following: a) the ability to ensure effective fixation in the intended location within the inferior vena cava; b) the ability to maintain adequate integrity; c) the ability to capture clots in the blood, while allowing acceptable blood flow; d) the ability to prevent clinical perforation/penetration; e) the compatibility of the filter dimensions for use with the specified caval diameters;
REQ-12 | 6.4 | shall | 6.4 Optional filter In addition to the attributes listed in 6.3, the design attributes of the optional filter shall at least take into account the following: a) the ability to be engaged; b) the ability to be retrieved; c) the ability to maintain adequate structural integrity through retrieval.
REQ-13 | 6.5 | shall | 6.5 Convertible filter In addition to the attributes listed in 6.3, the design attributes of the convertible filter shall at least take into account the following: a) the ability to be engaged; b) the ability to be converted (e.g. into a stent); c) the ability to maintain adequate structural integrity through and after conversion.
REQ-14 | 6.6 | shall | 6.6 Retrieval system The design attributes of the retrieval system shall at least take into account the following: a) the ability to safely reach the filter location; b) the ability to safely engage the filter; c) the ability to safely retrieve the filter; d) the ability to inject contrast via the retrieval system if indicated in the IFU; e) the ability to be safely withdrawn.
REQ-15 | 6.7 | shall | 6.7 Conversion system The design attributes of the conversion system shall at least take into account the following: a) the ability to safely reach the filter location; b) the ability to safely engage the conversion component of the filter; c) the ability to safely convert the filter; d) the ability to inject contrast via the conversion system if indicated in the IFU; e) the ability to be safely withdrawn.
REQ-16 | 6.8 | shall | 6.8 Filter system, retrieval system and conversion system In addition to the general requirements, the design attributes of the filter system, the retrieval system and the conversion system, as applicable, shall at least take into account the following: a) visibility under fluoroscopy or other applicable imaging modalities; b) compliance with the requirements of ISO 10993-1 and other appropriate parts of the ISO 10993 series; c) the sterility of the systems; d) the ability to minimize blood loss (haemostasis);
REQ-17 | 6.9 | shall | The design attributes of a coating on a delivery system or filter shall at least take into account the following:
REQ-18 | 6.10 | shall | The design attributes of an absorbable filter, and a filter containing an absorbable component or an absorbable coating, shall at least take into account the following:
REQ-19 | 8.1 | shall | The requirements of ISO 14630:2012, Clause 7, apply. A risk analysis shall be carried out in accordance with the requirements of ISO 14971.
REQ-20 | 8.1 | shall | Because optional filters can be used as permanent filters, testing appropriate for a permanent filter shall be conducted for optional filters. Additional testing is appropriate for optional filters.
REQ-21 | 8.1 | shall | If the convertible filter can be used as a permanent filter, testing appropriate for a permanent filter system shall be conducted. If part of the convertible filter is retrieved, testing appropriate for an optional filter system shall be conducted. If the convertible filter is intended to function as a stent after conversion, the requirements of ISO 25539-2 are applicable after conversion. If the convertible filter is not intended to function as a stent after conversion, additional testing addressing the risks associated with the remaining implant can be appropriate.
REQ-22 | 8.1 | shall | The device design concept shall be considered in the selection of appropriate tests and associated test methods. The device design concept includes the:
REQ-23 | 8.1 | shall | A device evaluation strategy shall be created. A device evaluation strategy provides the rationale for the testing selected to evaluate the filter system, retrieval system or conversion system, based on the requirements of the device design and potential failure modes. The device evaluation strategy may be communicated in a table (device evaluation strategy table) as described in Annex A. Alternative methods for presenting the device evaluation strategy may be used (e.g. a non-tabular presentation of the rationale for the testing based on the potential risks and benefits of the various systems for the intended clinical use).
REQ-24 | 8.1 | shall | Whenever changes are made in materials, construction, configuration, intended clinical use or processing methods, an appropriate analysis of the potential impact of the change on the potential failure modes and performance of the system shall be performed. This evaluation may be communicated using appropriate tables as described in Clause A.2. Appropriate testing shall be conducted as deemed necessary, considering the potential impact on device performance of the change.
REQ-25 | 8.1 | shall | Testing to establish the labelled shelf-life shall be conducted by repeating appropriate tests after aged conditions, where applicable. Justification for the selection of tests shall be provided. Generally, this will not include long-term durability testing, unless the materials of construction are susceptible to degradation that cannot be evaluated through shorter-term testing, or other tests that measure parameters that are not expected to be affected by aging (e.g. MRI safety testing, corrosion testing).
REQ-26 | 8.2 | shall | A sampling plan shall be utilized which ensures adequate representation of the design has been obtained for each characteristic measured. It should be verified that the design attributes of the filter system, retrieval system or conversion system are representative of the devices to be released for distribution including all sizes, configurations and components.
REQ-27 | 8.2 | shall | The sampling shall fully represent the range of device sizes and does not necessarily require testing of each size. It can be necessary to conduct an analysis to identify the size(s) of the device with the greatest potential for failure.
REQ-28 | 8.2 | shall | The samples selected for each test shall represent adequately challenging case(s). Consideration shall be given to filter size and orientation, delivery system sizes (diameter and length) and implant conditions (e.g. intended vena cava size and shape). It can be necessary to conduct an analysis to identify the samples with the greatest potential for failure under specified implant conditions.
REQ-29 | 8.3 | shall | All samples shall be subjected to sterilization, including multiple sterilizations, if appropriate, unless justification is provided for use of non-sterilized products.
REQ-30 | 8.4 | shall | — sampling: state the sampling plan, including the basis for and the number of samples tested; selection of test articles shall be justified (e.g. sizes, conditioning);
REQ-31 | 8.4 | shall | — acceptance criteria, if applicable: state the criteria for the test results, including justification and/or clinical relevance; clinical applicability of the acceptance criteria shall take into consideration the anatomical and physiological conditions of the intended use;
REQ-32 | 8.5.1 | shall | Testing of the filter system, the filter, the filter retrieval system and the filter conversion system shall be conducted to evaluate the design attributes in Clause 6, as applicable. The appropriate tests to evaluate each design attribute are based on the potential associated failure modes, device effects of failure and clinical effects of failure. The rationale for the requirements specified in this document for the bench tests and analytical analyses to assess device performance is described in Annex A.
REQ-33 | 8.5.2.1 | shall | The ability of the filter system to permit safe and consistent delivery, deployment and withdrawal shall be assessed.
REQ-34 | 8.5.2.1 | shall | Testing shall include the items listed in 8.5.2.2 to 8.5.2.6, as appropriate to the design of the filter system.
REQ-35 | 8.5.2.4.2 | shall | Evaluate the ability of the filter system to be advanced through the introducer sheath. Torquability shall be
REQ-36 | 8.5.2.6 | shall | filter system separately or concurrently with the torsional bond strength evaluation. The results shall be
REQ-37 | 8.5.3.1 | shall | The ability of the implant to filter blood for clots without causing caval occlusion shall be assessed.
REQ-38 | 8.5.3.1 | shall | Testing shall include the items listed in 8.5.3 to 8.5.5, as appropriate to the design of the filter. The tests are grouped based on similarities in the objectives of the testing; however, tests are not repeated within multiple categories. Refer to Annex A for a complete listing of the test applicable to each design attribute.
REQ-39 | 8.5.3.2 | shall | Determine the ability of the filter to capture clots, demonstrating that the device can capture clinically significant emboli, yet still permit sufficient blood flow around trapped emboli without caval occlusion. Clinically significant emboli, usually described as the numbers of clots at various sizes (e.g. diameters and lengths), shall be determined and justified by the manufacturer. The clot-trapping ability of the filter shall also be challenged in non-optimal positions (e.g. tilted) to determine the sensitivity of the design. Refer to References [48], [52], [73], [75], [80], [82] to [86] and [104] to [106] for more information on clot trapping.
REQ-40 | 8.5.3.3 | shall | In cases where different metals can be in contact by virtue of the device design or IFU, galvanic corrosion shall be assessed (e.g. radiopaque markers). Corrosion assessment includes, but is not limited to, the evaluation of test results, the review of literature and the consideration of the historical clinical performance of the material(s) under assessment. Guidance on corrosion assessment is given in a variety of sources (e.g. literature, textbooks, standards, regulatory guidance documents).
REQ-41 | 8.5.3.5.1 | shall | Potential integrity failures to be assessed can include fracture and component separation. Relevant in vivo loads (e.g. compression resulting from respiration and Valsalva manoeuvre) shall be considered when evaluating fatigue and durability.
REQ-42 | 8.5.3.5.1 | shall | In vitro fatigue testing of the filter or the appropriately justified test article shall be performed to demonstrate durability for an appropriate and justified duration. For example, in vitro testing representing a minimum of 10 years should be conducted for permanent implants. If the intended implant duration is less than 10 years, then shorter duration fatigue testing is more appropriate, with justification.
REQ-43 | 8.5.3.5.1 | shall | If fatigue testing is performed to compare the durability of one filter to another filter that has clinically demonstrated durability or clinically known problems with durability, the duration of the test shall be justified.
REQ-44 | 8.5.3.5.1 | shall | Consideration shall be given specifically to convertible filters in both filtering and converted states.
REQ-45 | 8.5.4.1 | shall | In addition to the attributes listed in 8.5.3.1, the ability of the optional filter to be safely, consistently and accurately retrieved shall be assessed.
REQ-46 | 8.5.4.1 | shall | Testing shall include the items listed in 8.5.4.2 to 8.5.4.4, as appropriate to the design of the optional filter.
REQ-47 | 8.5.5.1 | shall | In addition to the attributes listed in 8.5.2.1, the ability of the convertible filter to be safely, consistently and accurately converted shall be assessed, including the use of the conversion system as specified in the IFU, as applicable. The associated device/procedure related functions, potential failure modes and potential device and clinical effects of failure to be considered are listed in Table A.6.
REQ-48 | 8.5.5.1 | shall | Testing shall include the items listed in 8.5.4.2 to 8.5.4.4, as appropriate to the design of the convertible filter.
REQ-49 | 8.5.6.1 | shall | The ability of the retrieval system to permit safe and consistent retrieval of the filter shall be assessed.
REQ-50 | 8.5.6.1 | shall | Testing shall include the items listed in 8.5.6.2 to 8.5.6.6, as appropriate to the design of the retrieval system.
REQ-51 | 8.5.6.5 | shall | Evaluate the torsional strength of the joints and/or fixed connections in the segments of the retrieval system that are subjected to torsion during clinical use. Evaluate the torsional strength of the segments adjacent to the bonds of the retrieval system separately or concurrently with the torsional bond strength evaluation. The results shall be evaluated in relation to the torque necessary to access, retrieve and withdraw the system. Alternatively, provide justification if torsion is not applicable based on the intended clinical use.
REQ-52 | 8.5.7.1 | shall | The ability of the conversion system to permit safe and consistent conversion of the filter shall be assessed.
REQ-53 | 8.5.7.1 | shall | Testing shall include the items listed in 8.5.6.2 to 8.5.6.6, as appropriate to the design of the conversion system.
REQ-54 | 8.5.7.5 | shall | Evaluate the torsional strength of the joints and/or fixed connections in the segments of the conversion system that are subjected to torsion during clinical use. Evaluate the torsional strength of the segments adjacent to the bonds of the conversion system separately or concurrently with the torsional bond strength evaluation. The results shall be evaluated in relation to the torque necessary to access, convert and withdraw the system. Alternatively, provide justification if torsion is not applicable based on the intended clinical use.
REQ-55 | 8.5.8.1 | shall | Testing shall include the items listed in 8.5.7.1 and 8.5.7.5, as appropriate to the design of all device components.
REQ-56 | 8.5.8.3 | shall | The biocompatibility of the filter system, the retrieval system and the conversion systems shall be evaluated in accordance with ISO 10993-1 and appropriate other parts of the ISO 10993 series. The filter should generally be tested separately from the delivery system. However, there can be some tests (e.g. thrombogenicity, implantation) where the filter and delivery system are used or tested together for at least part of the test, and these should be appropriately justified.
REQ-57 | 8.5.8.4 | shall | Sterilization shall be ensured in accordance with appropriate international standards.
REQ-58 | 8.5.8.6.1 | shall | Document the risk assessment. If the risk assessment indicates that either acute and/or chronic particulate generation does not require further evaluation based on the level of risk, the risk assessment documentation is sufficient to address the particulate generation requirement. If the risk assessment indicates that the potential for acute and/or chronic particulate generation and likely clinical significance exists, acute and/or chronic testing shall be conducted, as applicable. It is not expected that particulates will always need to be characterized beyond size and quantity. Depending on the quantified particulate test results and associated clinical risk, additional particulate characterization (e.g. chemical identification, crystallinity, morphology) and identification of the particulate source can be appropriate.
REQ-59 | 8.5.9 | shall | Based on the potential associated failure modes, device effects of failure and clinical effects of failure, identify the appropriate testing to evaluate the specific design attributes associated with a coating on a delivery system or filter and conduct testing to address the risks identified in the assessment. Testing shall be identified as part of the device evaluation strategy and shall include an assessment of shelf-life.
REQ-60 | 8.5.10 | shall | Based on the potential associated failure modes, the device effects of failure and the clinical effects of failure, identify the appropriate testing to evaluate the specific design attributes associated with an absorbable filter or a filter containing an absorbable coating and conduct testing to address the risks identified in the assessment. Testing shall be identified as part of the device evaluation strategy and shall include an assessment of shelf-life.
REQ-61 | 8.6.1 | should | The purpose of preclinical in vivo testing is to evaluate the deployment of the filter, the biological response of the host to the filter, the effect of the implant environment on the filter, and the retrieval or conversion of the filter. Preclinical in vivo testing should provide data pertaining to safety. If the objective of an animal study can be met through alternative means (e.g. through reference to previously conducted animal and/or clinical studies), the use of previously obtained data or other supportive information shall be justified. The justification should include comment on the relevance of any differences between the subject device and the device used in the previous study and the relevance of any differences in the intended uses. Multiple studies may be conducted to address all the relevant specific aims for a particular filter system, a retrieval system and/or a conversion system.
REQ-62 | 8.6.2 | shall | Specific aims of the study shall be stated in the protocol. More than one study may be used to address these aims which can include the following:
REQ-63 | 8.6.3 | shall | Each filter system, retrieval system and/or conversion system shall be tested by implantation of the filter at the intended, or an analogous, vena cava site in a reasonable number of animals for an adequate duration of time (e.g. 12 weeks) to accomplish the specific aims of the study. A control can be appropriate for comparison purposes. For retrievability, an appropriate duration prior to filter retrieval shall be justified. Similarly, for conversion studies, an appropriate duration prior to conversion shall be justified. The type and intervals of interim assessments shall be specified and justified. For novel technologies, interim sacrifices and longer implant durations may be indicated. As far as permitted by the limitations of the animal model, all devices used should be of clinical quality and size and of the design intended for clinical use.
REQ-64 | 8.6.3 | shall | All animals in the study shall be regularly examined. Histological and pathological assessment of explants and appropriate tissues/organs shall be completed. If an animal either dies or must be sacrificed prior to scheduled termination, it shall be subjected to immediate post-mortem examination. The cause of death or illness, and the extent to which the implant was implicated shall be documented. Information for all animals implanted with either test or control filter, including those excluded from the final analyses, shall be recorded and included in the test report.
REQ-65 | 8.6.3 | shall | The design of the preclinical in vivo testing, including the experimental protocol, measurement methods and data analysis, shall be documented. In addition, the choice of animal model, such as species, gender, age, weight, shall be justified and shall be consistent with the study objectives. Implantation, and as applicable retrieval or conversion, shall be consistent with the instructions for use, as far as permitted by the limitations of the animal model. In particular, the filter sizing with respect to the IVC should be as close to the indicated sizing as possible in order to realistically simulate implantation conditions.
REQ-66 | 8.6.4 | shall | The following minimum data shall be recorded for each animal receiving a filter:
REQ-67 | 8.6.5 | shall | 8.6.5 Test report and additional information Results of all animals enrolled in the protocol shall be recorded and reported, even if excluded from the final analysis.
REQ-68 | 8.6.5 | shall | The test report shall include the following: a) the study protocol; b) the rationale for selection of the following: 1) animal model, including species, gender, age and weight; 2) implantation site; 3) the control for comparison or rationale for not using a control group (if applicable); 4) implantation periods; 5) methods of assessment; 6) intervals of observation;
REQ-69 | 8.7.1 | shall | The purpose of clinical evaluation is to assess the safety and effectiveness of the filter system, the retrieval system and/or the conversion system. An investigation shall be carried out for each new filter or new clinical application of a filter using the principles given in ISO 14155 or an equivalent reference. Significant design changes that can impact safety and performance shall require clinical evaluation if determined to be necessary based on an appropriate risk assessment. This evaluation may follow the requirements described in this document or the abbreviated requirements as appropriate to address the identified risks. Additional filter sizes outside the previously evaluated range can require clinical evaluation but does not always require assessment consistent with all requirements (e.g. multicenter study, statistically powered sample size).
REQ-70 | 8.7.1 | shall | If an objective of a clinical study can be met through alternative means (e.g. through reference to previously conducted clinical studies), the use of previously obtained data or other supportive information shall be justified. The justification should include comment on the relevance of any differences between the subject device and the device used in the previous study and the relevance of any differences in the intended uses.
REQ-71 | 8.7.2 | shall | Specific aims of the study shall be based on an appropriate risk assessment for the filter system, the retrieval system and/or the conversion system stated in the protocol. The specific aims may include the following:
REQ-72 | 8.7.3 | shall | A multicenter study shall be performed at a minimum of three investigational sites. A justification for the number of investigational sites shall be provided.
REQ-73 | 8.7.3 | shall | A specific question or set of questions (i.e. hypotheses) shall be defined prospectively. These questions shall delineate the appropriate safety (e.g. freedom from major adverse events), effectiveness (e.g. technical success in absence of serious device related events) or combined safety and effectiveness end points to be measured. Definitions of success and failure for each end point and the duration of follow-up needed to assess each end point shall be specified. A definition for the study success shall also be specified (e.g. meeting both the safety and effectiveness primary end points).
REQ-74 | 8.7.3 | shall | A statistical justification for the number of patients studied shall be provided based upon the primary hypotheses. Preferably no investigational site should enroll more than 35 % of the total number of study subjects.
REQ-75 | 8.7.3 | shall | The duration of the patient follow-up shall be determined in relation to the objectives of the clinical investigation. Patient follow-up intervals shall include a minimum of an assessment at discharge and at the specified study duration. A justification will be required for follow-up intervals. All patients enrolled in the study, including those excluded from the primary end point analyses, shall be recorded and reported. The final report may be completed when the required number of patients to test the hypotheses have reached the specified study duration. The report shall include current follow-up data on all patients. Longer-term patient follow-up (e.g. 2 years to 5 years after the filter has been implanted) can be appropriate for the post-market clinical assessment of device designs with a limited history of clinical use.
REQ-76 | 8.7.3 | shall | A method for evaluating the clinical outcomes shall be prospectively defined and justified (e.g. performance goals or a justifiable control group).
REQ-77 | 8.7.3 | shall | The study design shall be designated by appropriate terms [e.g. number of study arms, type of control (randomized, literature, performance goal), blinding, prospective versus retrospective]. The assessments can differ between the control group and the treatment group. If so, the assessments for both the control and treatment subjects shall be included in the protocol.
REQ-78 | 8.7.3 | shall | Patient inclusion and exclusion criteria shall be clearly identified. The criteria shall specify the target population (i.e. those for whom the filter is intended) and the accessible population (i.e. those you agree and are able to participate fully in the study). An appropriate epidemiological approach shall be utilized for recruiting subjects to minimize bias (e.g. encourage sequential enrolment).
REQ-79 | 8.7.3 | shall | Definitions, primary and secondary clinical end points, measurement methods and data analysis shall be specified in the clinical protocol. Secondary end points can include the following: — individual components that make up any composite primary end points; — change in quality of life status or other relevant patient reported outcomes; — measure of therapeutic success; — technical success (e.g. successful deployment of filter at intended implantation site); — procedural success (e.g. technical success in absence of serious device-related adverse events at 30 days);— retrieval or conversion success; — device and clinical effects of failure; — longer-term outcomes (e.g. 12-month safety data if the primary safety end point is at 30 days).
REQ-80 | 8.7.4 | shall | 8.7.4 Data acquisition Data similar to those for patients in the treatment group shall be recorded for each patient in the control group, as appropriate.
REQ-81 | 8.7.4 | shall | At a minimum, the following data shall be recorded for each patient in the treatment group: a) identification and demographic data: 1) patient identification; 2) indication for treatment (e.g. evidence of PE or DVT in the IVC or lower extremity veins and either an absolute or relative contraindication to anticoagulation, a failure of anticoagulation, or high risk from complication to anticoagulation) and associated medical diagnosis (e.g. lower extremity DVT); 3) demographics: i) age; ii) sex; iii) weight or body mass index;
REQ-82 | 8.7.5 | shall | 8.7.5 Final report The final report shall include the following: a) study protocol, including at a minimum: 1) study description (e.g. study design designation, control arm, number of sites, number of patients); 2) primary and secondary end points, hypotheses and definitions of success; 3) source of the data (e.g. site, core lab); 4) definition of study success; 5) subject population (i.e. selection criteria); 6) follow-up intervals; 7) methods of assessment (e.g. clinical, venography, duplex ultrasound); 8) data analysis plan including methods to address missing data; 9) definitions of technical and procedural success, device and clinical effects of failure, and adverse events; b) rationale, based on the risk assessment and questions to be answered, for selection of the following: 1) study size; 2) choice of control; 3) measurement methods; 4) statistical analyses employed; 5) patient follow-up intervals; c) number of patients treated at each investigational site; d) follow-up accountability (e.g. number of patients eligible for each follow-up interval and the number with specified follow-up data), including rationale for the exclusion of data from the primary end point analyses; e) demographics and risk factors; f) significant and/or relevant deviations from the clinical protocol and the manner in which deviations were addressed in the data presentation; g) results: 1) technical success; 2) procedural success; 3) retrieval or conversion success;
REQ-83 | 8.7.5 | shall | 9 Post-market experience A systematic procedure to review post-market experience gained from implants shall be in place using the principles given in ISO 14630:2012, 7.4, ISO 14155, ISO/TR 20416 and ISO 14971, or equivalent publications.
REQ-84 | 8.7.5 | shall | 10 Manufacturing Filter systems, retrieval systems and conversion systems shall be manufactured in such a way that the design attributes are achieved. Requirements are specified in other related international standards. The requirements of ISO 13485 and ISO 14630:2012, Clause 8 shall apply.
REQ-85 | 11.1 | shall | 11.1 Products supplied sterile Filter systems, retrieval systems and conversion systems that are labelled “sterile” shall comply with international, national or regional standards and have a sterility assurance level (SAL) of 10−6. Sterilization processes shall be validated and routinely controlled.
REQ-86 | 11.1 | shall | a) For systems that are to be sterilized by ethylene oxide, ISO 11135 shall apply.
REQ-87 | 11.1 | shall | b) For systems that are to be sterilized by moist heat, ISO 17665-1 shall apply.
REQ-88 | 11.1 | shall | c) For systems that are to be sterilized by radiation, the ISO 11137 series shall apply.
REQ-89 | 11.1 | shall | d) For systems that are to be sterilized by other sterilization processes, ISO 14937 shall apply.
REQ-90 | 11.2 | shall | 11.2 Sterilization residuals The requirements of ISO 14630:2012, 9.4 shall apply.
REQ-91 | 12.1.1 | shall | The requirements of ISO 14630:2012, Clause 10 shall apply.
REQ-92 | 12.1.2 | shall | Each filter system, retrieval system and conversion system shall be packaged in a unit container providing a sterile barrier, if applicable. It shall be readily apparent if the unit container has been opened.
REQ-93 | 12.1.3 | shall | Each unit container shall be packaged in an outer container. This outer container shall be designed to protect the unit container from damage caused by storage.
REQ-94 | 12.1.4 | shall | Each outer container, or a number of outer containers not necessarily of the same type, shall be packaged in a shipping container designed to protect the contents under normal conditions of handling, transit and storage.
REQ-95 | 12.1.5 | shall | For filter systems, retrieval systems and conversion systems supplied sterile, the unit container shall be designed to maintain the sterility of the system under normal conditions of handling, transit and storage, and to permit the contents to be presented for use in an aseptic manner.
REQ-96 | 12.1.5 | shall | The packaging shall conform to ISO 11607-1.
REQ-97 | 12.2.1 | shall | Filter system, retrieval system and conversion system shall be accompanied by a label(s) on an appropriate container(s).
REQ-98 | 12.2.2 | shall | At least the following information shall be provided in words, phrases, symbols or drawings on the label(s):
REQ-99 | 12.2.3 | shall | 12.2.3 Retrieval systems At least the following information shall be provided in words, phrases, symbols or drawings on the label(s): a) applicable information as described in 12.2.2; b) retrieval system information, including at least: 1) dimensions, i.e. the size of introducer (internal diameter), the required size of guidewire and the usable, working or effective length of the catheter; 2) recommended accessory devices; 3) retrieval system orientation.
REQ-100 | 12.2.4 | shall | 12.2.4 Conversion systems At least the following information shall be provided in words, phrases, symbols or drawings on the label(s): a) applicable information as described in 12.2.2; b) conversion system information, including at least: 1) dimensions, i.e. the size of introducer (internal diameter), the required size of guidewire and the usable, working or effective length of the catheter; 2) recommended accessory devices; 3) conversion system orientation.
REQ-101 | 12.3.1 | shall | 12.3.1 General The requirements of ISO 14630:2012, 11.3 shall apply.
REQ-102 | 12.3.2 | shall | 12.3.2 Information and instructions for use Each unit container or outer container of which the contents are identical shall be supplied with the IFU of the device or instructions on how to access an electronic version of the IFU. The instructions shall include the following information to use the device safely and properly, taking into account the training and knowledge of the potential users: a) name, address and/or trademark of the manufacturer; b) product name; c) device description and materials of construction; 1) material of construction of the filter; 2) description of the location(s) of markers for visualization, if applicable; d) indications for use and as applicable intended use; e) contraindications, cautions and warnings; f) recommendations for filter sizing, including vena cava diameters, as applicable; g) potential adverse events; h) data from clinical studies, if applicable; i) recommended methods for the aseptic presentation and the preparation of the filter and delivery system; j) recommended methods for vessel preparations, such as pre-dilatation, and methods for access, delivery of the filter, retrieval of the filter, conversion of the filter and withdrawal of the system; k) applicable information regarding the sterility of the device, in prominent form: 1) NON-STERILE; 2) STERILE; 3) STERILE – DO NOT RESTERILIZE; l) statement that the device is single use only in prominent form; m) resterilization information, if applicable; n) notification of additives or leachable components (if applicable); o) recommendations for storage, if applicable; p) date of or reference relating to the publication of the IFU, indicating when the IFU was last revised; q) recommendations for visualization; r) MRI safety information; s) recommended method, where appropriate for retrieval or conversion of the filter; for convertible filters, include the time to conversion; for retrievable filters, include the statement(s) that vena cava filter
REQ-103 | 12.3.2 | shall | A sampling plan shall be utilized which ensures that adequate representation of the design has been obtained for each characteristic measured. It should be verified that the design attributes of the filter system, the retrieval system or the conversion system are representative of the devices to be released for distribution including all sizes, configurations and components.
REQ-104 | 12.3.2 | shall | The sampling shall fully represent the range of device sizes and does not necessarily require testing of each size. It can be necessary to conduct an analysis to identify the size(s) of the device with the greatest potential for failure.
REQ-105 | 12.3.2 | shall | The samples selected for each test shall represent the adequately challenging case(s). Consideration shall be given to filter size and orientation, delivery system sizes (diameter and length) and implant conditions (e.g. intended vena cava size and shape). It can be necessary to conduct an analysis to identify the samples with the greatest potential for failure under specified implant conditions.
REQ-106 | 12.3.2 | shall | Clinical applicability of the acceptance criteria shall take into consideration the anatomical and physiological conditions of the intended use: — test method: describe in detail the method used to perform the test, including any prospectively defined inspection procedures, and provide a justification for relevant test parameters; — protocol deviations: describe any deviations and their potential significance on the interpretation of the results; — expression of results: report testing results expressed in units as indicated in the test method;— discussion, if applicable: include a discussion on the potential clinical significance of the results;— conclusions: state conclusions based on the comparison of results to acceptance criteria or provide an explanation of the relevance of the results for characterization tests.
REQ-107 | 12.3.2 | shall | D.5.2.1.3 Sampling Sampling shall be in accordance with Clause D.2.
REQ-108 | 12.3.2 | shall | D.5.2.1.4 Conditioning Conditioning shall be in accordance with Clause D.3.
REQ-109 | 12.3.2 | shall | D.5.2.1.6 Expression of results Length shall be expressed in centimetres (cm). Other dimensions shall be expressed in millimetres (mm). Results regarding the compatibility of the recommended accessory devices and the verification of the lumen and outer diameters, if applicable, should be documented.
REQ-110 | 12.3.2 | shall | The test report shall be in accordance with Clause D.4. The test report shall include the maximum, minimum, mean and standard deviation of all measured dimensions, the results of any verified dimensions, and the results of the observations of the accessory compatibility.
REQ-111 | 12.3.2 | shall | An assessment of the parameters that affect the force to deploy a particular system design shall be considered in designing an appropriate anatomical model. Literature and patient data are appropriate sources to identify challenging anatomy. Selection of the model material and model geometry should take into consideration the compliance of the vasculature being represented by the model. The expected response of the in vivo vessel to the insertion of accessory devices (e.g. guidewire, introducer sheath) and the filter system, and the friction associated with the model material, should also be considered when selecting the model material and test fluid for:
REQ-112 | 12.3.2 | shall | Sampling shall be in accordance with Clause D.2. Filter systems to be tested should be representative of the devices that have the potential for the highest deployment force.
REQ-113 | 12.3.2 | shall | Conditioning shall be in accordance with Clause D.3.
REQ-114 | 12.3.2 | shall | The test report shall be in accordance with Clause D.4 and shall include the maximum, minimum, mean and standard deviation of the deployment forces, and any anomalous observations. The report shall include a description of and justification for the anatomical model (e.g. angulation, tortuosity, diameter and construction of material of the model).
REQ-115 | 12.3.2 | shall | Sampling shall be in accordance with Clause D.2.
REQ-116 | 12.3.2 | shall | Conditioning shall be in accordance with Clause D.3.
REQ-117 | 12.3.2 | shall | 1) Evaluate the ability of the filter system to be advanced to, and positioned in, the targeted deployment location of the anatomical model(s) without compromising the function of the delivery system. This testing shall include, during the advancement of the system, the ability to torque the system without negatively affecting the ability to deploy. Note any anomalies and their impact on the performance of the filter system.
REQ-118 | 12.3.2 | shall | The results of the evaluation in the simulated use test shall be expressed descriptively. Distances shall be expressed in millimetres (mm).
REQ-119 | 12.3.2 | shall | The test report shall be in accordance with Clause D.4 and shall include all results and abnormal observations. The report shall include a description of the anatomical model(s) used and justification of how the model(s) is representative of the anatomical variation in the intended patient population (i.e. angulation, tortuosity and diameter). The test fluid and the model material of construction shall be reported and justified. The results for advancement, withdrawal (e.g. pushability, flexibility, trackability, torquability, as applicable)
REQ-120 | 12.3.2 | shall | and compatibility between the accessory devices and the filter system shall each be reported. Additionally, observed anomalies identified during deployment, the conformability of the deployed filter to the vessel wall, filter positioning (including orientation, if applicable) and observed filter anomalies shall be reported. Deployment accuracy shall be reported taking into account directionality, including the maximum, minimum, mean and standard deviation. The type and location of any filter or delivery system damage and any clinically relevant accessory device damage shall be reported.
REQ-121 | 12.3.2 | shall | The purpose of this test is to determine the tensile strength of all joints and/or fixed connections of the filter system (e.g. distal tip, hub bond). The strength of the segments adjacent to the bonds of the filter system shall be evaluated separately or concurrently with the bond strength determination.
REQ-122 | 12.3.2 | shall | Sampling shall be in accordance with Clause D.2.
REQ-123 | 12.3.2 | shall | Conditioning shall be in accordance with Clause D.3. Conditioning of the test samples should include loading, tracking (access and withdrawal) and deployment. Multiple tracking cycles through an appropriate anatomical model should be considered. Information regarding an appropriate anatomical model is provided in D.5.2.3.2. Filter systems from completed simulated use testing (see D.5.2.3) may be used for this test.
REQ-124 | 12.3.2 | shall | Bond strength shall be expressed in newtons (N).
REQ-125 | 12.3.2 | shall | The test report shall be in accordance with Clause D.4 and shall include the type and location of the failure and the maximum, minimum, mean and standard deviation of the bond strength(s).
REQ-126 | 12.3.2 | shall | The purpose of this test is to evaluate the torsional strength of the joints and/or fixed connections in the segments of the filter system that are subjected to torsion during clinical use. The torsional strength of the segments adjacent to the bonds of the filter system shall be evaluated separately or concurrently with the torsional bond strength evaluation.
REQ-127 | 12.3.2 | shall | Sampling shall be in accordance with Clause D.2.
REQ-128 | 12.3.2 | shall | Conditioning shall be in accordance with Clause D.3. Conditioning of the test samples should include loading, tracking (access and withdrawal) and deployment. Multiple tracking cycles through an appropriate anatomical model should be considered. Information regarding an appropriate anatomical model is provided in D.5.2.3.2. Filter systems from completed simulated use testing (see D.5.2.3) may be used for this test.
REQ-129 | 12.3.2 | shall | c) Attach the other end of the test sample to the torque gauge. The test gauge length shall be long enough to avoid influence of the clamping apparatus and short enough to ensure uniform torsional loading.
REQ-130 | 12.3.2 | shall | D.5.2.5.6 Expression of results Torsional bond strength (torque strength) shall be expressed in newton-metres (N·m) or number of rotations.
REQ-131 | 12.3.2 | shall | D.5.2.5.7 Test report The test report shall be in accordance with Clause D.4 and shall include the mode and location of the failure and the maximum, minimum, mean and standard deviation of the torsional bond strength.
REQ-132 | 12.3.2 | shall | The results shall be evaluated in relation to the torque level necessary to access, deploy and withdraw the system.
REQ-133 | 12.3.2 | shall | Sampling shall be in accordance with Clause D.2.
REQ-134 | 12.3.2 | shall | Conditioning shall be in accordance with Clause D.3.
REQ-135 | 12.3.2 | shall | The test report shall be in accordance with Clause D.4 and shall include the minimum, maximum, mean and standard deviation of the clot-trapping ability and the pressure gradient for each condition tested. The number of associated clots, device configuration, model configuration and clot size should also be reported. The anatomical model, including diameter(s) and length(s), clot material, test fluid and flow rate(s) used during testing should also be reported. Any anomalous observations should be reported as well.
REQ-136 | 12.3.2 | shall | — filter or appropriate test samples of the filter (e.g. segments, sections, components, subassemblies) that have undergone actual or simulated manufacturing processes; the test sample shall be appropriate to the type of corrosion under evaluation (e.g. crevice, pitting, fretting, galvanic);
REQ-137 | 12.3.2 | shall | Sampling shall be in accordance with Clause D.2.
REQ-138 | 12.3.2 | shall | Conditioning shall be in accordance with Clause D.3 and shall include actual or simulated loading and deployment.
REQ-139 | 12.3.2 | shall | Test data shall be expressed in units appropriate to the methods selected.
REQ-140 | 12.3.2 | shall | The test report shall be in accordance with Clause D.4 and shall include the complete corrosion assessment, including a summary of all test data, analyses and referenced information, comparisons to applicable controls, any appropriate comparison between in vivo and in vitro performance and conclusions regarding the anticipated corrosion resistance of the filter. For quantitative data, the maximum, minimum, mean and standard deviation shall be included. Applicable requirements indicated in documents used for testing should also be included.
REQ-141 | 12.3.2 | shall | The analyses shall be performed on the sizes and configurations necessary to ensure an adequate evaluation of the filter.
REQ-142 | 12.3.2 | should | 2) Establish the filter geometry and, if appropriate, mock or diseased vessel geometry. The geometry should be representative of the finished product. Consideration shall be given to the allowed variability of the dimensions when selecting the geometry for analysis.
REQ-143 | 12.3.2 | shall | D.5.3.3.4 Expression of results Stress shall be expressed in megapascals (MPa). Strain shall be expressed as a percentage (%) or dimensionless. Locations of critical stresses and/or strains should be depicted in colour figures with legends.
REQ-144 | 12.3.2 | shall | Diagrams for fatigue analysis shall be provided (e.g. constant life, Morrow analysis, Goodman analyses).
REQ-145 | 12.3.2 | shall | The report shall include the following: a) background and purpose: 1) provide a brief device description and the intended use environment;
REQ-146 | 12.3.2 | shall | Sampling shall be in accordance with Clause D.2.
REQ-147 | 12.3.2 | shall | Sampling shall allow for the evaluation of the structural integrity of all relevant parts of the filter.
REQ-148 | 12.3.2 | shall | The filter size(s) and configurations with the greatest potential for fatigue failure shall be identified and justified using computational modelling (e.g. finite element analysis), engineering analysis and/or clinical data. Alternatively, samples representing the range of sizes may be chosen for testing.
REQ-149 | 12.3.2 | shall | Conditioning shall be in accordance with Clause D.3.
REQ-150 | 12.3.2 | shall | 3) test frequency; the test frequency shall be selected to maintain the test article deformation within the pre-defined limits for the duration of the test and to avoid undesirable harmonics, localized heating of the filter and rate-dependent effects on material properties;
REQ-151 | 12.3.2 | shall | c) testing: after initiation of testing, at periodic intervals, monitor the test conditions and equipment operation to ensure that the test article does not migrate and experiences the intended displacements or forces; because the relationship between the intended displacements or forces and test article deformation can change over time, it can be necessary to verify that the test article is deforming as intended; the methodology to verify that the displacements and/or deformations are as intended shall be described in the test method; if appropriate, stop the test at periodic intervals for inspection of the test article(s); removing the test article from a mock vessel for the periodic inspection is not recommended. However, if removing the test article from a mock vessel is necessary, care shall be taken to remove and re-deploy in a manner that minimizes the effect on the test article;
REQ-152 | 12.3.2 | should | The test report should done be according to Clause D.4. The intended and measured displacements and/or deformations shall be reported.
REQ-153 | 12.3.2 | shall | Results of all inspections, including the cycle count at which the inspections took place and the number and location of any observed anomalies, including fractures, shall be reported. The test report should include a discussion on the potential causes (e.g. fatigue failure, material inclusion, pre-existing sample damage, mock
REQ-154 | 12.3.2 | shall | The test frequency shall be expressed in cycles per second (Hz). Displacements shall be expressed in millimetres (mm). Compressive deformation shall be expressed as a percentage (%).
REQ-155 | 12.3.2 | shall | Refer to the information in the general test method, test report subclause (see D.5.3.4.1.7) with the following: the intended and measured minimum and maximum test article compressive displacement shall be reported.
REQ-156 | 12.3.2 | shall | Sampling shall be in accordance with Clause D.2.
REQ-157 | 12.3.2 | shall | Conditioning shall be in accordance with Clause D.3 and shall include loading, preconditioning and deployment.
REQ-158 | 12.3.2 | shall | Outer diameter(s) shall be measured at appropriate locations after deployment in accordance with the IFU. For non-circular cross-sections, it can be appropriate to measure and report the maximum and minimum values.
REQ-159 | 12.3.2 | shall | Diameters shall be expressed in millimetres (mm). Lengths shall be expressed in millimetres (mm) or centimetres (cm).
REQ-160 | 12.3.2 | shall | The test report shall be in accordance with Clause D.4 and shall include the maximum, minimum, mean and standard deviation of all measured and calculated dimensions.
REQ-161 | 12.3.2 | shall | Sampling shall be in accordance with Clause D.2.
REQ-162 | 12.3.2 | shall | Conditioning shall be in accordance with Clause D.3.
REQ-163 | 12.3.2 | shall | The test report shall be in accordance with Clause D.4 and shall include the type and location of the failure, and the maximum, minimum, mean and standard deviation of tensile strength.
REQ-164 | 12.3.2 | shall | D.5.3.7.3 Sampling Sampling shall be in accordance with Clause D.2.
REQ-165 | 12.3.2 | shall | D.5.3.7.4 Conditioning Conditioning shall be in accordance with Clause D.3 and shall include loading and preconditioning.
REQ-166 | 12.3.2 | shall | The test report shall be in accordance with Clause D.4 and shall report the presence of migration and include the maximum, minimum, mean and standard deviation of the peak pressure differential across the filter, or peak force, for each condition tested. The anatomical model, including diameter(s) and material should be reported. For the migration pressure gradient measurement, the clot material, the clot size(s), the test fluid, the flow rate(s) and pressure used during testing should also be reported.
REQ-167 | 12.3.2 | shall | Sampling shall be in accordance with Clause D.2.
REQ-168 | 12.3.2 | shall | Conditioning shall be in accordance with Clause D.3 and shall include loading and preconditioning.
REQ-169 | 12.3.2 | shall | The test report shall be in accordance with Clause D.4 and should include a description of the filter feature(s) tested (e.g. fixation legs), the minimum, maximum, mean and standard deviation of the radial force at the minimum and maximum diameters for each device size tested. Results from expansion should be reported, with the respective speeds used during testing.
REQ-170 | 12.3.2 | shall | D.5.6.2.3 Sampling Sampling shall be in accordance with Clause D.2.
REQ-171 | 12.3.2 | shall | D.5.6.2.4 Conditioning Conditioning shall be in accordance with Clause D.3.
REQ-172 | 12.3.2 | shall | D.5.6.2.7 Test report The test report shall be in accordance with Clause D.4 and shall include the maximum, minimum, mean and standard deviation of the peak force and observations. The report shall include a description of and justification for the anatomical model (e.g. angulation, tortuosity, diameter and construction of material of the model).
REQ-173 | 12.3.2 | shall | Sampling shall be in accordance with Clause D.2.
REQ-174 | 12.3.2 | shall | Conditioning shall be in accordance with Clause D.3.
REQ-175 | 12.3.2 | shall | The results of the evaluation in the simulated use test shall be expressed descriptively.
REQ-176 | 12.3.2 | shall | The test report shall be in accordance with Clause D.4 and shall include all results and abnormal observations. The report shall include a description of the anatomical model(s) used and justification of how the model(s) is representative of the anatomical variation in the intended patient population (i.e. angulation, tortuosity and diameter). The test fluid and the model material of construction shall be reported and justified. The applicable results for pushability, flexibility, torquability, trackability and compatibility between the accessory devices and the retrieval system shall each be reported. Additionally, observed anomalies identified during retrieval shall be reported. The type and location of any retrieval system damage and any clinically relevant accessory device damage shall be reported.
REQ-177 | 12.3.2 | shall | D.5.6.4 Torsional bond strength The purpose of the torsional bond strength test is to evaluate the torsional strength of the joints and/or fixed connections in the segments of the retrieval system that are subjected to torsion during clinical use. The torsional strength of the segments adjacent to the bonds of the retrieval system shall be evaluated separately or concurrently with the torsional bond strength evaluation.
REQ-178 | 12.3.2 | shall | D.5.6.5 Tensile bond strength The purpose of the tensile bond strength test is to determine the bond strength of all joints and/or fixed connections of the retrieval system (e.g. distal tip, hub bond). The strength of the segments adjacent to the bonds of the retrieval system shall be evaluated separately or concurrently with the bond strength determination.
REQ-179 | 12.3.2 | shall | D.5.7.2.3 Sampling Sampling shall be in accordance with Clause D.2.
REQ-180 | 12.3.2 | shall | D.5.7.2.4 Conditioning Conditioning shall be in accordance with Clause D.3.
REQ-181 | 12.3.2 | shall | D.5.7.2.7 Test report The test report shall be in accordance with Clause D.4 and shall include the maximum, minimum, mean and standard deviation of the peak force and observations. The report shall include a description of and justification for the anatomical model (e.g. angulation, tortuosity, diameter and construction of material of the model).
REQ-182 | 12.3.2 | shall | D.5.7.3.3 Sampling Sampling shall be in accordance with Clause D.2.
REQ-183 | 12.3.2 | shall | D.5.7.3.4 Conditioning Conditioning shall be in accordance with Clause D.3.
REQ-184 | 12.3.2 | shall | b) Following the IFU and using the appropriate accessory devices (e.g. guidewire, introducer sheath), deploy the filter in the model, insert the conversion system, convert the filter and withdraw the system while evaluating the following applicable attributes: 1) the ability to access the conversion location in the anatomical model: this testing shall include the evaluation of pushability, flexibility, trackability and torquability as applicable; note any anomalies and their impact on the performance of the conversion system; 2) the ability to deliver the conversion system to the filter location; 3) the ability to engage the filter with the system; 4) the ability to convert the filter; 5) the ability to withdraw the system.
REQ-185 | 12.3.2 | shall | The results of the evaluation in the simulated use test shall be expressed descriptively.
REQ-186 | 12.3.2 | shall | The test report shall be in accordance with Clause D.4 and shall include all results and abnormal observations. The report shall include a description of the anatomical model(s) used and justification of how the model(s) is representative of the anatomical variation in the intended patient population (i.e. angulation, tortuosity and diameter). The test fluid and the model material of construction shall be reported and justified. The results for compatibility between the accessory devices, including pushability, flexibility, trackability, torquability, as applicable, and the retrieval/conversion system shall each be reported. Additionally, observed anomalies identified during conversion shall be reported. The type and location of any conversion system damage and any clinically relevant accessory device damage shall be reported.
REQ-187 | 12.3.2 | shall | The purpose of this test is to evaluate the torsional strength of the joints and/or fixed connections in the segments of the conversion system that are subjected to torsion during clinical use. The torsional strength of the segments adjacent to the bonds of the conversion system shall be evaluated separately or concurrently with the torsional bond strength evaluation.
REQ-188 | 12.3.2 | shall | The purpose of this test is to determine the bond strength of all joints and/or fixed connections of the conversion system (e.g. distal tip, hub bond). The strength of the segments adjacent to the bonds of the conversion system shall be evaluated separately or concurrently with the bond strength determination.
REQ-189 | 12.3.2 | shall | Sampling shall be in accordance with Clause D.2.
REQ-190 | 12.3.2 | shall | Conditioning shall be in accordance with Clause D.3.
REQ-191 | 12.3.2 | shall | The results of the assessments shall be expressed descriptively, with representative images as appropriate.
REQ-192 | 12.3.2 | shall | The test report shall be in accordance with Clause D.4 and shall include the assessment of visibility for all applicable components at the various stages of testing. Describe the results of the assessments and/or include visual results (e.g. representative fluoroscopic images). The test report shall also include the make and model of the imaging equipment, the relevant imaging parameters and details of the phantom tissue model.
REQ-193 | 12.3.2 | shall | Sampling shall be in accordance with Clause D.2.
REQ-194 | 12.3.2 | shall | Conditioning shall be in accordance with Clause D.3.
REQ-195 | 12.3.2 | shall | Particle size shall be expressed in micrometres (μm) and should be presented in appropriate size bins (e.g. ≥10 μm, ≥25 μm, ≥50 μm).
REQ-196 | 12.3.2 | shall | The test report shall be in accordance with Clause D.4. The maximum, minimum, mean and standard deviation of the particle counts in each size bin shall be reported.